ELECTRON BEAM LOSS IN COMMERCIAL ESEM

- Gerasimos D Danilatos1, Matthew R Phillips2, John V Nailon3

1ESEM Research Laboratory, 98 Brighton Boulevard, North Bondi, NSW 2026, Australia.

2Microstructural Analysis Unit, University of Technology, Sydney, PO Box 123, Broadway, NSW 2007, Australia.

3Centre for Microscopy and Microanalysis, The University of Queensland, Brisbane, Qld 4072, Australia.

Key words: environmental scanning electron microscope, electron scattering, electron loss, differential pumping, gas jet, pressure limiting aperture, image noise, probe loss.

Short title:  Beam Loss in Commercial ESEM

Address for correspondence:
GD Danilatos, 98 Brighton Boulevard, North Bondi, NSW 2026, Australia.

Telephone:  +61 2 91302837

Fax:  +61 2 93650326

Email:  esem@bigpond.com

ELECTRON BEAM LOSS IN COMMERCIAL ESEM

- Gerasimos D Danilatos1, Matthew R Phillips2, John V Nailon3

1ESEM Research Laboratory, 98 Brighton Boulevard, North Bondi, NSW 2026, Australia

2Microstructural Analysis Unit, University of Technology, Sydney, PO Box 123, Broadway, NSW 2007, Australia

3Centre for Microscopy and Microanalysis, The University of Queensland, Brisbane, Qld 4072, Australia

# Abstract

The electron beam in an environmental scanning electron microscope suffers small but significant losses even before it enters the specimen chamber of the instrument.  The losses are due to the gaseous jet formed inside and downstream of the pressure limiting aperture, they increase with specimen chamber pressure and eventually become catastrophic.  These effects are minimal in an optimum design system.  Two commercial environmental scanning electron microscopes have been tested by the use of the direct simulation Monte Carlo method.  The electron loss has been calculated based on the mass thickness computed for the gas.  It has been found that the losses are well above the level expected for an optimum design and clearly above the level originally experienced with an experimental prototype microscope.

## Introduction

Conventional differential pumping is used on present day commercial environmental scanning electron microscope (ESEM) instruments to separate the high vacuum of the electron optics column from the high pressure specimen chamber.  A transition region of intermediate pressures is established between the two extreme regions of the microscope.  The electron beam is transferred from the high vacuum to the high pressure via a transition region in which the beam undergoes collisions with the gas and suffers initial electron losses.  The loss of electrons in the transition region is an inevitable consequence of the differential pumping method used.  In particular, the magnitude of electron loss is determined by the design of a given instrument and can be greater or equal to some physical limit imposed by the formation of a supersonic gas jet downstream of the pressure-limiting aperture (PLA).

An early experimental prototype of ESEM was designed and operated close to the physical limit of minimum electron beam loss in the transition region (Danilatos, 1981).  The optimum condition was experimentally determined to occur when (a) the wall thickness of the PLA was much less than its diameter and (b) the conductance of the downstream evacuation pipe was as large as possible, which, together with sufficient pump speed, created a minimum back-pressure downstream of the transition region.  The relatively thin-walled PLA was obtained by the use of commercially available copper grid apertures of different diameters.  The reported thickness of PLA wall was 10 μm and 30 μm.  A detailed experimental study of the behavior of such apertures including the formation of a supersonic jet was later reported by Danilatos (1983); all subsequent work by the same author has applied the same principles (Danilatos, 1988).

The consequence of correctly applying differential pumping together with appropriate detection techniques has allowed ESEM to operate at any pressure between vacuum and one atmosphere.  Different applications require different specimen chamber pressures, which can be divided in various utilitarian (functional, or practical) pressure ranges, such as:  (a) Pressure sufficient to suppress charging effects, usually up to around 200 Pa; (b) pressure sufficient to maintain liquid water phase, i.e. greater than 609 Pa; (c) pressure corresponding to the saturation water vapor pressure at room temperature, i.e. around 2 kPa; (d) saturation water vapor pressure at “living body” temperature (36 0C), i.e. around 6 kPa; (e) gas pressure that sustains animal life, i.e. around 20-100 kPa; (f) an open-ended ESEM, i.e. an instrument without a specimen chamber, capable of operating at full atmospheric pressure for in-situ examination of large objects.

Two commercial ESEM instruments have been evaluated in regard to gas density variation and electron beam loss in the transition region of their differential pumping systems, and a comparison is made to the prototype ESEM operating close to the optimum condition.  This study has been prompted by concerns that a significant deviation from expected standards has been found in practice.  This deviation consists in lower signal-to-noise-ratio at a given pressure and in lower upper pressure limit of operation at a given accelerating voltage, for each instrument.  It is shown that an unnecessary deterioration of performance is due, by and large, to a departure from the optimum geometry design of the apertures.

## Materials and Methods

In order to find the properties of the gas flow through a PLA, an alternative to the experimental techniques employed in past work is to theoretically calculate them.  However, the calculation was not always possible or practical, because analytical formulations of the gas flow exist only for the extreme cases of either continuum flow (high pressures/small apertures) or free-molecule flow (low pressure/large apertures).  The conditions of ESEM operation correspond to the transition region between continuum and free-molecule flow.  Even where analytical solutions are applicable, the optimum ESEM geometry corresponds to very difficult or impractical cases to solve.  Fortunately, the gas flow properties can be computed by simulation techniques such as the direct simulation by Monte Carlo (DSMC) method (Bird,1985).

The DSMC method is a technique for the computer modeling of a real gas by some thousands, or even millions of simulated molecules depending on the gas pressure and the physical extent of the flow field.  The velocities and positions of these molecules are stored in the computer and are modified with time as the molecules move and collide through the boundaries defining the flow field.  The entry and exit conditions for a given gas are initially set and the program computes the gas properties until it reaches a steady-state situation, whereupon the program continues to sample the equilibrium properties until a satisfactory smooth average for each field point is achieved.  These programs were initially developed for space engineering problems involving large space vehicles in rarefied gas conditions, and required the use of mainframe computers.  The availability of fast and cheap personal computers, which can be devoted to solve DSMC problems, has allowed this technique to be adapted and used by the present author for gas flow computations in the ESEM (Danilatos, 1991); (Danilatos, 1993).  The same technique is used in the present work to compute three cases of PLA configurations and the associated gas density variation.

The first case involves a 0.5 mm diameter PLA on a 0.1 mm thick plate with the rim of the aperture tapered at 45 degrees angle diverging in the downstream direction of the gas flow.  This geometry of PLA together with the gas flow is shown in Fig. 1, where only half of the PLA cross-section is drawn on one side of the aperture axis, because the flow field is axially symmetric.  The specimen chamber of the ESEM is located on the left-hand side and the leaking gas through the aperture is pumped out from the right had side.  The electron beam travels along the axis in the direction from low pressure to high pressure.

The second case involves the PLA assembly of an ElectroScan E3 model ESEM employed by the University of Queensland (QLD) since 1991, the axi-symmetrical semi-cross-section of which is drawn in Fig. 2.  In the same way, the specimen chamber is located on the left-hand side of the drawing, whereby gas enters from side CR and exits from side IJ.  Some gas also exits from the second PLA formed on plate LM.  The electron beam travels along the axis RQ.

The third case involves the PLA assembly of a Philips XL30 ESEM employed by the University of Technology, Sydney (UTS), since 1999, the axi-symmetrical cross-section of which is drawn in Fig. 3.  Similarly, the specimen chamber is located on the left-hand side of the drawing, whereby gas enters from side CU and exits from side LM.  Some gas also exits from the second PLA formed on plate OP.  The electron beam travels along the axis UT.

The electron beam encounters gas molecules with which it collides well before it enters in the specimen chamber.  A gas plume or jet forms downstream of (or “above”) the PLA plane.  The gas density decreases rapidly along the axis of the jet but the precise variation of density depends on the geometry of walls within which the gas is constrained.  The fastest density decrease is expected to occur in the first case (Fig. 1) but a quantitative evaluation of this phenomenon is necessary in order to establish the actual consequences on imaging.  The latter is achieved by finding the amount of electron scattering in the plume by the use of electron scattering collision theory.

The electron beam scattering process is governed by the Poisson distribution probability P(x)

P(x)=mxe−mx!	(1)

which gives the probability for an electron to undergo x number of collisions, when the average number of collisions per electron is m.  Knowledge of the parameter m allows us to immediately find the electron beam current I that is transmitted completely without any scattering by the gas molecules, when the initial incident beam current in vacuum is I0.  The fraction of transmitted beam is given by the exponential equation

II0=e−m	(2)

The parameter m is found from the function of number density n(z)along the axis z and the total scattering cross-section σT of the gas:

m=σTn(z)dz	(3)

The above integration is performed between any two limits defining the gas layer of interest.  In the present situation, we integrate from the entrance interface of the PLA up to a point where the density is essentially reduced to a negligible value.  As upper limit of integration we take 10 mm for the first case corresponding to Fig. 1 (thin PLA), 13.7 mm for the second case (QLD) and 16 mm for the third case (UTS).  In the latter two cases, we use the location of the second PLA shown in Figs. 2 and 3 as the upper limit of integration.  The integral factor can be referred to as the molecular thickness (molthick) corresponding to the customary term of mass thickness (i.e. density times thickness) in other microscopy works, namely:

molthick=n(z)dz	(4)

This mathematical analysis is correct in absolute terms, but, in practice, the predicted values are reliable only to the extent that the supplied values of total scattering cross-section and density variation are correct.  The DSMC method has been shown to be reliable in practice so that the supplied value of the total scattering cross-section alone determines the accuracy of the electron beam loss found in this work.  Argon is used as the test gas, because its scattering cross-section can be derived analytically as a monatomic gas (Jost and Kessler, 1963); (Danilatos, 1988); the values used are tabulated for different accelerating voltages below.  However, one of the main aims of this work is to produce the gas density variation, from which the molecular thickness is found.  The latter can be used alone to evaluate instrument performance, as it is a measure of the total amount of gas in the electron beam path “above” the PLA.

The specimen chamber pressure and the pressure at the exit boundaries of the system are set to a desired value and the computer finds the density at each point inside the flow field.  Because there is a depletion zone immediately upstream of (i.e. “below”) the PLA plane, the entry gas boundary is set 2 mm further upstream, where the density is essentially equal to the stagnation density in the specimen chamber held at room temperature, i.e. 293 K.  Along with the density, the computer also finds the temperature, velocity and Mach number at every point of the flow field, as well as the leak rate of gas through various interfaces, such as through the PLA.  These and other characteristics will be included in a more comprehensive report to be released in due course, later.

## Results

Two specimen chamber pressures are considered, namely, 100 Pa and 1000 Pa, which are typical of commercial ESEM operation at present.  As far as the first case of Fig. 1 is concerned, the back-pressure at the pump is set to 0 Pa corresponding to an ideal physical limiting case.  We consider four accelerating voltages and the results are given in Table 1.

The back-pressure at the exit interfaces is not exactly known for the given instruments but four typical values are chosen for the interface where the differential pump acts (referred to as pump pressure).  These four pump pressures are expected to cover a range within which the real situation is expected to fall.  These values are 0, 1, 2 and 4 Pa when the specimen chamber pressure is 100 Pa, and 0, 10, 20 and 40 Pa when the specimen chamber pressure is 1000 Pa.  The pressure at the second PLA interface with the rest of the electron optics column is assumed to be zero in all cases (high vacuum interface).  Therefore, there are eight combinations of pressure for each commercial instrument, at four accelerating voltages each combination, which are given in Table 2 for the QLD instrument and in Table 3 for the UTS instrument.

Graphical representations of some typical cases are also given below.  Fig. 4 shows the variation of gas particle density (atoms per cubic meter) along the axis of the three systems, when the specimen chamber is held at 100 Pa and the pump pressure is 0 Pa.  Each curve represents the average density of four radial positions along the axis, namely, at 0.05, 0.10, 0.15 and 0.20 mm, where the density varies slightly.

Fig. 5 shows the density variation along the jet axis for the three systems when the specimen chamber pressure is 1000 Pa and the pump pressure is 0 Pa.

Fig 6 shows the variation of gas density along the axis of the QLD system, when the specimen chamber is held at 100 Pa and the pump pressure is 0, 1, 2 and 4 Pa.  For comparison, the ideal case curve (thin aperture at 0 Pa pump pressure) is also included.

Fig 7 shows the variation of gas density along the axis of the QLD system, when the specimen chamber is held at 1000 Pa and the pump pressure is 0, 10, 20 and 40 Pa.  For comparison, the ideal case curve (thin aperture at 0 Pa pump pressure) is also included.

Fig 8 shows the variation of gas density along the axis of the UTS system, when the specimen chamber is held at 1000 Pa and the pump pressure is 0, 10, 20 and 40 Pa.  For comparison, the ideal case curve (thin aperture at 0 Pa pump pressure) is also included.

Finally, a graphical representation of the transmitted electron fraction is shown in Fig. 9 for some typical cases: The fraction is given for all three systems with specimen chamber either at 100 or 1000 Pa, with 0 Pa pump pressure.  The QLD and UTS systems are also graphed with 20 Pa pump pressure.

## Discussion

The variation of density and molecular thickness for argon “above” the PLA has been established for an optimum (i.e. “thin”) and two commercial ESEM systems.  Based on the molecular thickness found, the fraction of electron beam that is totally unaffected in its path through the PLA system has been calculated for four accelerating beam voltages.  The electron beam continues to suffer even greater losses as it traverses a gas layer in the specimen chamber before it strikes a specimen, the details of which have been presented previously (Danilatos, 1993).  Here, a comparative study of the behavior of different systems, which can be characterized primarily by the gas above the PLA plane, has been presented. The effects in the neighborhood immediately below the PLA are generally of a second order in the high pressure regime.  The latter effects together with the effect of the positioning of a specimen below the PLA will be the subject of separate reports dealing with optimum design and imaging conditions of the differential pumping system of an ESEM.

From Table 1, we find that in all cases there is sufficient electron beam probe left for imaging even for an accelerating voltage as low as 5 keV, at 100 Pa of argon in the specimen chamber.  An even better situation should be expected with gases having a little lower scattering cross-section such as water vapor and nitrogen, since the present author has reported good imaging at pressures greater than 1000 Pa at 5 keV.

From Table 2, we find that both commercial instruments perform well at 100 Pa specimen chamber pressure but serious problems arise as soon as the pressure is raised to 1000 Pa.  In particular, the QLD instrument would have a problem operating at 5 keV and 1000 Pa, but it could just cope by increasing the accelerating voltage above 15 or 20 keV, provided that the exit pressure at the pump is maintained very low.  If the pump pressure is found to be more than, say, 20 Pa, then a significant component of molecular thickness is added as shown in the Table and in Fig. 7.  The area under each curve in the corresponding figures represents the molecular thickness; the deviation from the ideal case is clear.

Similarly, from Table 3 we find a similar but much more accentuated trend for the UTS instrument.  In fact, the electron beam losses are dramatic as soon as we reach the 1000 Pa pressure in specimen chamber.  The difference from the QLD and the ideal case becomes apparent in Figs. 4 and 5, also by comparing Figs. 7 and 8.

A comparative presentation is also shown in Fig. 9 for all three systems at their best possible situation, namely, with 0 Pa pump pressure.  In particular, the QLD and UTS systems are also shown with pump pressure at 20 Pa.  Operation seems impossible with both instruments at 5 keV and 1000 Pa.

The actual fraction of electron beam scattered (or transmitted) requires measurement of the actual pump pressure.  From the four values of pump pressure computed herewith, the corresponding parameters can be interpolated for a given pump pressure.  The latter is expected to lie somewhere in the range of pump pressures presented.

To find the effect of electron beam loss with other gases we need to repeat the present study.  However, recalling Avogadro’s law, which states that equal numbers of different gas particles occupy the same volume at the same temperature and pressure, we may assume that the molecular thickness for different gases is roughly the same.  The latter assumption is not strictly correct, because we are dealing with a dynamic situation of gas flow, whereas Avogadro’s law applies to static gases, so that the DSMC computation should be repeated for each gas separately.  However, the main difference with different gases arises from the difference in their scattering cross-section, so that, to a first approximation, we may get an idea of the electron beam losses by applying different scattering cross-sections to the molecular thickness tabulated herewith.

One remedy to overcome the electron beam loss problem is use very high accelerating voltage, but this is not desirable in most applications.  Especially with organic and insulating specimens, use of high accelerating voltages results in large specimen interaction volumes (or beam penetration), which usually results in specimen damage, instability, and loss of resolution.  The effect of large interaction volume is particularly pronounced in the backscattered electron mode of detection.

Another remedy to these problems is to use a smaller diameter PLA, which reduces the over all molecular thickness.  However, this also reduces the field of view at low magnifications, which may be undesirable for many applications.

The best solution when using conventional differential pumping system is to redesign the ESEM instrument to allow incorporation of an optimum PLA configuration.  Apart from a thin plate PLA, a conical geometry has also been studied by Danilatos (1993).  The determination of optimum thickness, cone angle, distance between the two pressure limiting apertures and the interplay of these parameters will be the subject of further work.

However, it is possible to overcome even the natural limit posed by the gas jet if we use the novel method of a reverse flow PLA (RF-PLA) announced recently (Danilatos, 2000).  According to the RF-PLA, an annular supersonic gas jet is introduced in the opposite direction around the PLA with a pumping action at its core.  By such mechanism, the conventional gas jet “above” the PLA is eliminated.  As a result, the electron beam is effectively free to travel and without loss above the PLA.

Another consideration is the study of distribution of electrons lost (or removed) from the electron beam into what has been termed “electron skirt”.  This is particularly important in x-ray microanalysis.  The study of electron skirt can be done experimentally, theoretically and computationally.  All hitherto studies assume an abrupt, or step-wise function of the gas density involved, i.e. the electron beam encounters no gas up to the PLA plane, after which it travels through a uniform gas layer inside the specimen chamber.  However, the present study has shown that the gas has a significant density “above” the PLA, which results in a significant electron loss.  In other words, a significant electron skirt has already formed prior to the beam entering the specimen chamber.  Any future computation of electron skirt incident on a specimen surface has to account also for the gas density variation both in the depletion zone “below” and in the gas jet “above” the PLA.  The DSMC method has provided the density at every point in the entire gas flow field, which can be used as input to any future calculations of electron skirt distributions.

## Conclusion

The conventional differential pumping system is characterized by a supersonic gas jet and gas plume formed downstream of a PLA.  This gas represents a mass-thickness, which the electron beam of an ESEM must overcome, before it enters in the specimen chamber. The mass-thickness results in certain amount of electron beam loss.  The gas density variation in the jet has been computed by the DSMC method.  These computations have been made for a thin plate PLA, which represents a natural limit of mass thickness. Two commercial ESEM instruments have also been studied and found to incur much greater losses.  As a consequence, these instruments cannot operate with an accelerating voltage of 5 keV, at 1000 Pa pressure of argon in the specimen chamber, while an optimum PLA design normally allows these conditions.  Quantitative studies to determine optimum conventional differential pumping systems can be done with the DSMC method.

References

Bird, GA (1985)  Molecular Gas Dynamics and the Direct Simulation of Gas Flows. Oxford: Oxford Science Publications, Clarendon Press.

Danilatos GD (1981)  Design and construction of an atmospheric or environmental SEM (Part 1). Scanning 4:9-20.

Danilatos GD (1983)  Design and construction of an atmospheric or environmental SEM-2.  Micron 14:41-52.

Danilatos GD (1988)  Foundations of environmental scanning electron microscopy.  Advances Electronics Electron Physics 71:109-250.

Danilatos GD (1991)  Gas flow properties in the environmental SEM. Microbeam Analysis-1991 (Ed. Howitt DG), San Francisco Press, San Francisco, pp. 201-203.

Danilatos GD (1993)  Bibliography of environmental scanning electron microscopy. Microsc. Res. Technique 25:529-534.

Danilatos GD (1993)  Environmental scanning electron microscope: Some critical issues. Scanning Microscopy Supplement 7:57-80

Danilatos GD (1999)  Reverse flow pressure limiting aperture (RF-PLA).  Microsc. Microanal. 6:21-30, 2000.

Jost K and Kessler J (1963)  Die Ortsverteilung mittelschneller Elektronen bei Mehrfachstreung, Zeits. Phys. 176:126-142.

Tables

| Table 1. THIN PLA |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| p, Pa | p1, Pa | molthick mols/m2 | E, eV | σT, m2 | m | II0 |
| 100 | 0 | 4.00E+18 | 5000 | 1.56E-20 | 0.06255 | 0.939 |
| 100 | 0 | 4.00E+18 | 10000 | 8.28E-21 | 0.03316 | 0.967 |
| 100 | 0 | 4.00E+18 | 15000 | 5.73E-21 | 0.02295 | 0.977 |
| 100 | 0 | 4.00E+18 | 20000 | 4.43E-21 | 0.01774 | 0.982 |
| 1000 | 0 | 5.69E+19 | 5000 | 1.56E-20 | 0.88874 | 0.411 |
| 1000 | 0 | 5.69E+19 | 10000 | 8.28E-21 | 0.47111 | 0.624 |
| 1000 | 0 | 5.69E+19 | 15000 | 5.73E-21 | 0.32602 | 0.722 |
| 1000 | 0 | 5.69E+19 | 20000 | 4.43E-21 | 0.25206 | 0.777 |

Table 1 Corresponding values of specimen chamber pressure p, pump pressure p1, molecular thickness molthick, accelerating voltage E, total scattering cross-section σT, average number of collisions m, transmitted fraction of beam II0 for a thin PLA.

| Table 2. UNIVERSITY OF QUEENSLAND ESEM (E3) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| p, Pa | p1, Pa | molthick mols/m2 | E, eV | σT, m2 | m | II0 |
| 100 | 0 | 9.70E+18 | 5000 | 1.56E-20 | 0.15144 | 0.859 |
| 100 | 0 | 9.70E+18 | 10000 | 8.28E-21 | 0.08028 | 0.923 |
| 100 | 0 | 9.70E+18 | 15000 | 5.73E-21 | 0.05555 | 0.946 |
| 100 | 0 | 9.70E+18 | 20000 | 4.43E-21 | 0.04295 | 0.958 |
| 100 | 1 | 1.31E+19 | 5000 | 1.56E-20 | 0.20479 | 0.815 |
| 100 | 1 | 1.31E+19 | 10000 | 8.28E-21 | 0.10856 | 0.897 |
| 100 | 1 | 1.31E+19 | 15000 | 5.73E-21 | 0.07512 | 0.928 |
| 100 | 1 | 1.31E+19 | 20000 | 4.43E-21 | 0.05808 | 0.944 |
| 100 | 2 | 1.61E+19 | 5000 | 1.56E-20 | 0.25125 | 0.778 |
| 100 | 2 | 1.61E+19 | 10000 | 8.28E-21 | 0.13318 | 0.875 |
| 100 | 2 | 1.61E+19 | 15000 | 5.73E-21 | 0.09217 | 0.912 |
| 100 | 2 | 1.61E+19 | 20000 | 4.43E-21 | 0.07126 | 0.931 |
| 100 | 4 | 2.23E+19 | 5000 | 1.56E-20 | 0.34792 | 0.706 |
| 100 | 4 | 2.23E+19 | 10000 | 8.28E-21 | 0.18443 | 0.832 |
| 100 | 4 | 2.23E+19 | 15000 | 5.73E-21 | 0.12763 | 0.880 |
| 100 | 4 | 2.23E+19 | 20000 | 4.43E-21 | 0.09867 | 0.906 |
| 1000 | 0 | 1.28E+20 | 5000 | 1.56E-20 | 1.99329 | 0.136 |
| 1000 | 0 | 1.28E+20 | 10000 | 8.28E-21 | 1.05662 | 0.348 |
| 1000 | 0 | 1.28E+20 | 15000 | 5.73E-21 | 0.73121 | 0.481 |
| 1000 | 0 | 1.28E+20 | 20000 | 4.43E-21 | 0.56532 | 0.568 |
| 1000 | 10 | 1.50E+20 | 5000 | 1.56E-20 | 2.33592 | 0.097 |
| 1000 | 10 | 1.50E+20 | 10000 | 8.28E-21 | 1.23824 | 0.290 |
| 1000 | 10 | 1.50E+20 | 15000 | 5.73E-21 | 0.85690 | 0.424 |
| 1000 | 10 | 1.50E+20 | 20000 | 4.43E-21 | 0.66249 | 0.516 |
| 1000 | 20 | 1.75E+20 | 5000 | 1.56E-20 | 2.73466 | 0.065 |
| 1000 | 20 | 1.75E+20 | 10000 | 8.28E-21 | 1.44962 | 0.235 |
| 1000 | 20 | 1.75E+20 | 15000 | 5.73E-21 | 1.00318 | 0.367 |
| 1000 | 20 | 1.75E+20 | 20000 | 4.43E-21 | 0.77558 | 0.460 |
| 1000 | 40 | 2.34E+20 | 5000 | 1.56E-20 | 3.65359 | 0.026 |
| 1000 | 40 | 2.34E+20 | 10000 | 8.28E-21 | 1.93673 | 0.144 |
| 1000 | 40 | 2.34E+20 | 15000 | 5.73E-21 | 1.34027 | 0.262 |
| 1000 | 40 | 2.34E+20 | 20000 | 4.43E-21 | 1.03620 | 0.355 |

Table 2.  Corresponding values of specimen chamber pressure p, pump pressure p1, molecular thickness molthick, accelerating voltage E, total scattering cross-section σT, average number of collisions m, transmitted fraction of beam II0 for the University of Queensland ESEM.

| Table 3. UNIVERSITY OF TECHNOLOGY, SYDNEY, ESEM (PHILIPS XL30) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| p, Pa | p1, Pa | molthick mols/m2 | E, eV | σT, m2 | m | II0 |
| 100 | 0 | 2.105E+19 | 5000 | 1.56E-20 | 0.32875 | 0.720 |
| 100 | 0 | 2.105E+19 | 10000 | 8.28E-21 | 0.17427 | 0.840 |
| 100 | 0 | 2.105E+19 | 15000 | 5.73E-21 | 0.12060 | 0.886 |
| 100 | 0 | 2.105E+19 | 20000 | 4.43E-21 | 0.09324 | 0.911 |
| 100 | 1 | 2.336E+19 | 5000 | 1.56E-20 | 0.36496 | 0.694 |
| 100 | 1 | 2.336E+19 | 10000 | 8.28E-21 | 0.19346 | 0.824 |
| 100 | 1 | 2.336E+19 | 15000 | 5.73E-21 | 0.13388 | 0.875 |
| 100 | 1 | 2.336E+19 | 20000 | 4.43E-21 | 0.10351 | 0.902 |
| 100 | 2 | 2.705E+19 | 5000 | 1.56E-20 | 0.42249 | 0.655 |
| 100 | 2 | 2.705E+19 | 10000 | 8.28E-21 | 0.22396 | 0.799 |
| 100 | 2 | 2.705E+19 | 15000 | 5.73E-21 | 0.15498 | 0.856 |
| 100 | 2 | 2.705E+19 | 20000 | 4.43E-21 | 0.11982 | 0.887 |
| 100 | 4 | 3.318E+19 | 5000 | 1.56E-20 | 0.51824 | 0.596 |
| 100 | 4 | 3.318E+19 | 10000 | 8.28E-21 | 0.27471 | 0.760 |
| 100 | 4 | 3.318E+19 | 15000 | 5.73E-21 | 0.19011 | 0.827 |
| 100 | 4 | 3.318E+19 | 20000 | 4.43E-21 | 0.14698 | 0.863 |
| 1000 | 0 | 2.691E+20 | 5000 | 1.56E-20 | 4.20410 | 0.015 |
| 1000 | 0 | 2.691E+20 | 10000 | 8.28E-21 | 2.22855 | 0.108 |
| 1000 | 0 | 2.691E+20 | 15000 | 5.73E-21 | 1.54222 | 0.214 |
| 1000 | 0 | 2.691E+20 | 20000 | 4.43E-21 | 1.19233 | 0.304 |
| 1000 | 10 | 2.888E+20 | 5000 | 1.56E-20 | 4.51102 | 0.011 |
| 1000 | 10 | 2.888E+20 | 10000 | 8.28E-21 | 2.39124 | 0.092 |
| 1000 | 10 | 2.888E+20 | 15000 | 5.73E-21 | 1.65481 | 0.191 |
| 1000 | 10 | 2.888E+20 | 20000 | 4.43E-21 | 1.27937 | 0.278 |
| 1000 | 20 | 3.113E+20 | 5000 | 1.56E-20 | 4.86233 | 0.008 |
| 1000 | 20 | 3.113E+20 | 10000 | 8.28E-21 | 2.57747 | 0.076 |
| 1000 | 20 | 3.113E+20 | 15000 | 5.73E-21 | 1.78368 | 0.168 |
| 1000 | 20 | 3.113E+20 | 20000 | 4.43E-21 | 1.37901 | 0.252 |
| 1000 | 40 | 3.724E+20 | 5000 | 1.56E-20 | 5.81746 | 0.003 |
| 1000 | 40 | 3.724E+20 | 10000 | 8.28E-21 | 3.08378 | 0.046 |
| 1000 | 40 | 3.724E+20 | 15000 | 5.73E-21 | 2.13406 | 0.118 |
| 1000 | 40 | 3.724E+20 | 20000 | 4.43E-21 | 1.64990 | 0.192 |

Table 3.  Corresponding values of specimen chamber pressure p, pump pressure p1, molecular thickness molthick, accelerating voltage E, total scattering cross-section σT, average number of collisions m, transmitted fraction of beam II0 for the University of Technology, Sydney, ESEM.

Figures

Figure 1.  Semi-cross-section of thin plate PLA with density contours of argon flowing from left (with specimen chamber pressure at100 Pa) to right (with pump at 0.00 Pa).

Figure 2. 2 Drawing of the geometry of the internal “bullet” assembly for the QLD) ESEM.  Line RQ is an axis of symmetry.  The point coordinates are A(0.6, 2.0), B(0.0, 0.8), C(0.0, 0.25), D(0.4, 0.25), E(1.0, 1.75), F(3.0, 1.75), G(3.0, 2.0), H(7.4, 2.0), I(7.4, 2.75), J(12.7, 2.75), K(12.7, 0.8), L(13.7, 0.8), M(13.7, 0.15), N(13.75, 0.15), O(13.75, 0.75), P(16.85, 0.75), Q(16.85, 0.0), R(0.0, 0.0).

Figure 3.  Drawing of the internal geometry of the "bullet" assembly of the ESEM at UTS.  Line UT is the axis of symmetry.  The point coordinates are A(1.0, 1.5), B(0.0, 0.8), C(0.0, 0.25), D(1.1, 0.25), E(1.1, 0.5), F(1.3, 1.1), G(3.1, 1.1), H(3.1, 4.1), I(4.6, 4.1), J(4.6, 1.42), K(9.3, 1.42), L(9.3, 2.5), M(15.0, 2.5), N(15.0, 1.0), O(16.0, 1.0), P(16.0, 0.15), Q(16.2, 0.15), R(16.2, 1.0), S(22.7, 1.0), T(22.7, 0.0), U(0.0, 0.0).

Figure 4.  Density variation of gas of the PLA plane along the axis for the THIN, QLD and UTS systems, starting at 100 Pa in the upstream specimen chamber.

Figure 5.  Density variation of gas downstream of the PLA plane along the axis for the THIN, QLD and UTS systems, starting at 1000 Pa in the upstream specimen chamber.

Figure 6.  Density variation of gas downstream of the PLA plane for the QLD system at four pump pressures, with 100 Pa in the specimen chamber.  The “thin” case is reproduced from Fig. 4.

Figure 7. Density variation of gas downstream of the PLA plane for the QLD system at four pump pressures, with 1000 Pa in the specimen chamber.  The “thin” case is reproduced from Fig. 5.

Figure 8.  Density variation of gas downstream of the PLA plane for the UTS system at four pump pressures, with 1000 Pa in the specimen chamber.  The “thin” case is reproduced from Fig. 5.

Figure 9.  Transmitted fraction of a 5 keV electron beam through the THIN, QLD and UTS systems of PLA at chamber/pump pressures of 100/0, 1000/0 and 1000/20 Pa.